TEMPLE UNIVERSITY

Office of the Vice President for Research

Research  Compliance

Institutional Biosafety Committee

STANDARD OPERATING PROCEDURE SOP# 3.0

BIOSAFETY LEVEL 2+ (BSL2+) PROCEDURES

Name of PI:_____________________________________________________________________________

Date of form completion: __________

List BSL2+ procedures: __________________________________________________________________

- Purpose

This standard operating procedure (SOP) describes the safe BSL2+ procedures.

- Approvals

All research using BSL2+ procedures must have IBC and IACUC (if working with animals) approvals prior to initiation of any research work.

- Trainings

Principal Investigator is responsible to train all researchers, staff, and animal care providers on all aspects of this SOP.

- Security

- Only the authorized personnel (researchers, animal care provider) can have access to the facilities performing BSL2+ procedures.

- Escort the service workers (facilities engineer or BSC certifier), regulatory (OSHA, PaDEP, DOT, USDA, AAALAC), EHRS, IACUC inspectors at all times to the BSL2+ room for services or inspection.

- Signage

- Post agent summary, biohazard signage, and the agent’s name on the door to the room when performing BSL2+ procedures. Room(s): __________________________________________

- Post a biohazard symbol and agent’s name to the incubators, -80C freezers, and animal cages when storing BSL2+ procedural agents. Room(s):__________________________________________

- Remove agent summary, biosafety signage and agent’s name from the door to the room when the agent is no longer present in the room.

- Safety Equipment

- Biosafety Cabinet (BSC)

    - Certify the BSC annually by outside vendor to ensure continued, proper operation. Room(s): ________________________________________

    - Place a beaker containing 10% bleach inside the BSC, soak the tips and pipettes used for BSL2+ procedures for 10 minutes to disinfect the agents.

    - Use a vacuum pump for liquid aspiration. Do not use house vacuum line system.

- Place a trap containing 10% bleach and an in-line HEPA filter to protect the vacuum pump from the backflow contamination of a BSL2+ agent.

- Place a sharps container near the BSC for collecting needles, syringes, tips, pipettes, and any sharps objects.

- Centrifuge

- Use sealable centrifuge cups.

- Load and unload the centrifuge cups in a BSC.

- Autoclave

- Place a biological indicator (BI) inside the autoclave along with the materials to be autoclaved for ensuring the effectiveness of the autoclave cycles once every week or once every 40 hours of operations. Records of the BI testing must be readily available.

- Emergency eyewash must be readily available.

- Check the eyewash station weekly. The facilities management check the eyewash station annually. Records of the weekly check must be readily available. Eyewash location: Room __________________________.

- Personal Protective Equipment (PPE)

- PPE are disposable gown, hair cover, mask, face shield, shoe covers, and gloves. PPE must be present at entrance of the areas working with BSL2+ procedures.

- PPE donning procedure

  - Don the disposable gown, hair cover, mask, face shield, shoe covers, and double gloves before entering the BSL2+ room.

Gown ties in back                                       Mask secures                                            Face shield

on head                                              secures on head

  - Cover open wounds with bandages before double-gloving. Extent gloves to overlap the gown so that the wrist area is completely covered.

Cover wounds with                                           Gloves extend

bandages                                                over gown cuffs

  - Gloves should frequently be wetted with disinfectant and should be changed when contaminated or torn.

- PPE doffin procedure

Before exiting the BSL2+ room

      - Doff the first layer of gloves.

Peel away from hand                  Slide ungloved finger          Create a bag from both gloves

Turing gloves inside-out             Peel off from insdie             Discard in a red bag

      - Doff the face shield, hair cover, mask, shoe covers.

Grasp head pieces                                                                 Untie the bottom, then top

Lift away from face and discard                                          Remove from face and discard

      - Doff the disposable gown

Unfasten ties                          Peel gown away from              Turn contaminated outside

neck and shoulder                   toward the inside and roll into a bundle

After exiting the BSL2+ room

    - Doff the inner gloves outside the BSL2+ room.

- Dispose All the PPE in a red bag.

Solid container required to have a lined red bag, a lid, and a biohazard symbol

- Wash hands with antimicrobial soap using a sink.

- Decontamination

Use 10% bleach to wipe down equipment and work surfaces for decontamination before and after performing BSL2+ procedures.

- Animal Microisolator Cage Changing Procedure (This section only applies for performing BSL2+ procedure in animal studies)

- Place the soiled microisolator cages in a BSC.  Room(s): _________________________________

- Soiled bedding

  - Empty soiled bedding material into double red biohazard bags within the BSC. Securely seal the bag and spray Quatricide®PV-15 solution to exterior of the outer bag before removal from BSC.

  - Place the red bags in the red bin located in the ULAR soiled side cage wash area.

  - Soiled bedding material will be picked up by a certified vendor and decontaminated off site.

- Soiled caging material and Hydropac™ water pouches

  - Place empty caging material (cages, diet trays, Lifespan™ platforms, filter tops) in double autoclave bags within the BSC. Securely seal the autoclave bags and spray Quatricide®PV-15 to the exterior of the outer bag before removal from the BSC.

  - Place used Hydropac™ water pouches and lixit valves in separate double autoclave bags within the BSC. Securly seal the autoclave bags. Label the outer bag as “Liquid Cycle”. Spray Quatricide®PV-15 to the exterior of the outer bag before removal from the BSC.

- Move all bags to be autoclaved to the ULAR soiled side cage wash for autoclaving.

- ULAR staff will autoclave the caging material. Processed autoclaved caging material will be sent through the cage washer and auotclaved water pouches and valves will be discarded.

ULAR will bag bedding/caging:   yes ___  no ___

Laboratory staff will bag bedding/caging: yes___ no ___

- Place the empty cages in double  autoclavable bags within a BSC. Securely seal the autoclave bag and spray 10% bleach solution to exterior of the outer bag before removal from BSC.

- Transport the bagged cages and beeding material to the soiled side cagewash.

- ULAR staff will autoclave the material, then process autoclaved caging material through the cage washer. Autoclaved bedding material will be picked up by a certified vendor.

ULAR will bag bedding/caging:   yes ___  no ___

Laboratory staff will bag bedding/caging: yes___ no ___

- Biological Waste Management

- Place all solid waste in a red bag and autoclave prior to be picked up by a certified vendor. Place all sharps items in a sharps container to be picked up by certified vendor. Chemically disinfect or autoclave all infectious liquid waste before release to sanitary sewer.

- Place animal carcasses in an autoclavable bag in a BSC, place the carcasses in the animal freezer in the BSL2+ room (Room: ________________________), autoclave the bagged animal carcasses before placing the autoclaved carcasses in the Animal Facility Carcass freezer (Room: __________________________).

- Spill Management

  - Small Spill

    - Don the disposable gown, mask, hair cover, shoe covers and face shield and double gloves.

    - Contain the small spill by covering with absorbent pads or paper towel.

    - Disinfect the spill by adding 10% bleach and 10 minutes contact time.

    - Carefully remove potentially contaminated protective clothing as mentioned in the section of PPE doffing procedure if there is a splash in the protective equipment.

    - Dispose all solid wastes generated from spill in a red bag. Place all sharps items must in a sharp container.

- Large Spill (any spill involving concentrated virus, more than 100 ml of virus container liquid, aerosol generated incident, and incident involving the centrifuge)

    - Notify individual in the room to evacuate.

    - Leave the room and notify PI and EHRS.

    - Carefully remove potentially contaminated protective clothing based on the PPE doffing procedure listed in section G if the PPE is contaminated.

    - Place the contaminated PPE in a red bag.

    - Do not enter the room until clearance by the EHRS.

    - EHRS will respond the large spill according to the following procedure.

- Allow the aerosolized particles from the large spill to settle for 2 hours.

- Post a signage of “DO NOT ENTER” should be posted at the entrance of BSL2+ procedural room.

    - Don the non-absorbent coveralls, hair cover, mask, shoe covers, face shield and double gloves.

    - Prepare a large volume of a 20% bleach solution.

    - Spray all surfaces and items with 20% bleach including ceilings, walls, bench surfaces, cabinets and floors. Allow 10 minutes contact time. Then, spray everything with 70% isopropyl alcohol.

    - Remove PPE and place in a red bag and seal the waste bag. Sharps items are disposed of in a sharps container.

- Emergency Response

- Wash the affected areas with copious amounts of lukewarm water in the event of personnel exposure such as eye, skin, mucous membranes or body fluids contact when performing a BSL2+ procedure.

- Report to PI and seek immediate medical evaluation if there is personnel exposure

    - During office hour

    - Health Science Center: Go to the Occupational Health Department, M-F, 8:00 AM to 5:00 PM.

    - Main Campus: Go to Employee Health Services, M-F, 8:30 am - noon; 1 pm - 4:40 pm.

    - After office hour

  - After 5 PM, go to Temple University Hospital Emergency Department.

  - Principal Investigator must immediately report the incident to the EHRS.

- Training record of this SOP

| Name (please print) | Signature | TU ID Number | Department | Position / Title |
| --- | --- | --- | --- | --- |